## Slide 1: HF-based wet etch tests
Etch rates on four different HF-based etchant was measured on SiOx (recipe: OXIDE_100nm_300C) and SiN-H2 free films (recipe: SiN_No_NH3_100nm_300C).
Etching performed on wafers supplied by LNL on 2025-10-13.
A low etch rate of only 0.02 nm/s (1.2 nm/min) in BOE (10:1) observed for the SiN film. This etch rate is unexpectedly low and more in line what would be expected by and LPCVD film.
A much higher BOE (10:1) etch rate observed for SiOx of 2.7 nm/s (160 nm/min), resulting in a selectivity to SiN of 138. This SiOx etch rate is reasonable based on numbers provided in literature and about half that of ALD-SiOx deposited using the Fiji tool.
Material | Etchant | Etch rate (nm/s)
SiOx | BOE (10:1) | 2.7E+00
SiOx | HF (3:100) | 1.7E+00
SiOx | HF (1:100) | 3.7E-01
SiOx | HF (1:1000) | 2.8E-02
SiNx - H2 free | BOE (10:1) | 2.0E-02
SiNx - H2 free | HF (3:100) | 2.3E-02
SiNx - H2 free | HF (1:100) | 8.2E-03
SiNx - H2 free | HF (1:1000) | 1.0E-03

## Slide 2: P251020_SPTS
P251020_SPTS_1
  Recipe: OXIDE_100nm_300C
  Samples:
    Fused Silica 2" wafer from WUS20500250XXXXXSNN1
    n-Si 2'' wafer
    Pyramid chip (E250424B_W1_R4C4)
    GaN-on-sapphire planar piece (cleaved from one X-60WJOB24040770-4 wafer)
  Ellipsometry
    Thickness: 96.24 nm
    n: 1.49
P251020_SPTS_2
  Recipe: SiN_No_NH3_100nm_300C
  Samples:
    Fused Silica 2" wafer from WUS20500250XXXXXSNN1
    n-Si 2'' wafer
    Pyramid chip (E250424B_W1_R4C3)
    GaN-on-sapphire planar piece(cleaved from one X-60WJOB24040770-4 wafer)
  Ellipsometry
    Thickness: 98.72 nm
    n: 2.02
Thickness non-uniformity seem much larger than what was observed on 200 mm wafers. The placement of wafers/chips on carrier wafer probably impacts gas flows or plasma densities, resulting in varying thicknesses.
P251020_SPTS_1
P251020_SPTS_2

## Slide 3: P251020_SPTS_1 (SiOx)
MLA lithography on LOR/S1805 resist stack followed by lift-off of 10 nm Ti + 200 nm contacts performed on pyramid chip.
Metal contacts fabricated on top of SiO2 film (anodes) and on large areas where scribing was performed down to n-GaN layer (cathodes).
  Average resistance between cathodes about 49 Ω.
Only planar n-GaN-on-sapphire sample measured electrically.
Three different anode sizes measured (314 µm2, 1257 µm2, 32668 µm2), two of each size.
Very similar J(V) characteristics between all devices and sizes.
  Exclusively vertical current through the SiO2-film.
Average breakdown electric field of 10.7 mV/cm.
Device | Breakdown voltage (V) | Breakdown field (MV/cm) | Anode area (µm2)
LEDS_Q1_G1 | 98 | 10.2 | 32668
LEDS_Q1_J1 | 103 | 10.7 | 32668
LEDS_Q1_L1_20A1 | 99 | 10.3 | 1257
LEDS_Q1_G1_10E1 | 111 | 11.5 | 314
LEDS_Q1_G1_10A1 | 111.5 | 11.6 | 314
LEDS_Q1_L1_20C1 | -94.5 | -9.8 | 1257

## Slide 4: P251020_SPTS_1 (SiN)
MLA lithography on LOR/S1805 resist stack followed by lift-off of 10 nm Ti + 200 nm contacts performed on pyramid chip.
Metal contacts fabricated on top of SiN film (anodes) and on large areas where scribing was performed down to n-GaN layer (cathodes).
  Average resistance between cathodes about 65 Ω.
Only planar n-GaN-on-sapphire sample measured electrically.
Three different anode sizes measured (314 µm2, 1257 µm2, 32668 µm2), two of each size.
Very similar J(V) characteristics between all devices and sizes.
  Exclusively vertical current through the SiN-film.
Average breakdown electric field of 11.1 mV/cm.
Device | Breakdown voltage (V) | Breakdown field (MV/cm) | Anode area (µm2)
LEDS_CR_B1_20C1 | 106 | 11.0 | 1257
LEDS_CR_A1_20A1 | 107.5 | 11.2 | 1257
LEDS_CR_C1_10E1 | 106 | 11.0 | 314
LEDS_CR_C1_10A1 | 109.5 | 11.4 | 314
LEDS_CR_A1 | 105.5 | 11.0 | 32668
LEDS_CR_B1 | 107 | 11.1 | 32668